## Reference Range
Parameter | Lower | Higher | Unit | Note | Entered Parameter | Unit | Allowable Range | Exception Symbol | Description
pH | 7.35 | 7.45 |  |  | Temperature (Temp) | OC | 15 - 45 |  | Outside Reference Range - High
pH (Ven0us) | 7.32 | 7.43 |  |  | Barometric Pressure (BP) | mmHg | 500 - 999 (default 760) |  | Outside Reference Range - Low
pCO2 (Male) | 35 | 48 | mmHg |  | Percent inspirred oxygen (FiO2) | % | 10 - 100 |  | Outside Critical Limit - High
pCO2 (Female) | 32 | 45 | mmHg |  | Mean Airway Pressure (MAP) | cm H2O | 0 - 999 |  | Outside Critical Limit - Low
pCO2 (Venous) | 26 | 52 | mmHg | -6 - 7 higher than arterial pCO2 |  |  |  |  | Outside Reportable Range - Greater Than
pO2 | 83 | 108 | mmHg | [5 empty] | Outside Reportable Range - Less Than
Na+ | 136 | 145 | mmol/L | [5 empty] | Result Incalculable
K+ | 3.5 | 5.1 | mmol/L | [5 empty] | Absorbance Error
Cl- | 98 | 107 | mmol/L | [5 empty] | Sulfohemoglobin Interference Detected
Ca++ | 1.15 | 1.33 | mmol/L | [5 empty] | High Turbidity Detected
Hct (Male) | 40 | 50 | % | [5 empty] | Interference Detected
Hct (Female) | 37 | 47 | % | [5 empty] | Micro Clot Detected
Gluc | 65 | 95 | mg/dL | [5 empty] | Temporary Sensor Error
Lac | 0.36 | 0.75 | mmol/L | [5 empty] | High Methemoglobin Warning
Lac (Vernous) | 0.56 | 1.39 | mmol/L | [5 empty] | Sufohemoglobin and High Methemoglobin Warning
tHb (Male) | 12.6 | 17.4 | g/dL | [5 empty] | Corrected for Sulfohemoglobin
tHb (Female) | 11.7 | 16.1 | g/dL | [5 empty] | iQM2 IntraSpect
O2Hb | 90 | 95 | %
COHb (Non Smokers) |  | 3 | %
COHb (Smokers) |  | 10 | %
MetHb |  | 1.5 | %
HHb | 1 | 5 | %
sO2 | 94 | 98 | %
TCO2 | 19 | 24 | mmol/L
TCO2 (Venous) | 22 | 26 | mmol/L
BE | -2 | 3 | mmol/L
HCO3- | 21 | 28 | mmol/L
HCO3- (Venous) | 22 | 29 | mmol/L
Anion Gap | 10 | 20 | mmol/L | (Na+ + K+) - (Cl- + HCO3-)

## Critical Limits
Parameter | Lower | Higher | Unit | Note | Entered Parameter | Unit | Allowable Range
pH | 7.2 | 7.6 |  |  | Temperature (Temp) | OC | 15 - 45
pH (Ven0us) |  |  |  |  | Barometric Pressure (BP) | mmHg | 500 - 999 (default 760)
pCO2 (Male) | 20 | 70 | mmHg |  | Percent inspirred oxygen (FiO2) | % | 10 - 100
pCO2 (Female) |  |  | mmHg |  | Mean Airway Pressure (MAP) | cm H2O | 0 - 999
pCO2 (Venous) |  |  | mmHg | higher than arterial pCO2
pO2 | 40 |  | mmHg
Na+ | 120 | 160 | mmol/L
K+ | 2.8 | 7.8 | mmol/L
Cl- | 80 | 120 | mmol/L
Ca++ | 0.75 | 1.6 | mmol/L
Hct (Male) | 18 | 60 | %
Hct (Female) |  |  | %
Gluc | 40 | 450 | mg/dL
Lac |  | 3.4 | mmol/L
Lac (Vernous) |  |  | mmol/L
tHb (Male) | 7 | 20 | g/dL
tHb (Female) |  |  | g/dL
O2Hb |  |  | %
COHb (Non Smokers) |  |  | %
COHb (Smokers) |  |  | %
MetHb |  |  | %
HHb |  |  | %
sO2 |  |  | %
TCO2 |  |  | mmol/L
TCO2 (Venous) |  |  | mmol/L
BE |  |  | mmol/L
HCO3- | 10 | 40 | mmol/L
HCO3- (Venous) |  |  | mmol/L
Anion Gap |  |  | mmol/L | (Na+ + K+) - (Cl- + HCO3-)

## Rumus
| sO2 |  |  | 100*[O2Hb/(O2Hb+HHb)] | %
| sO2 (c) |  |  | (23400*(pO2^3+150*pO2)^-1+1)^-1 | %
| C (K) |  |  | (pO2/26.7)^0.184
| BE(B) Siggaard Anderson |  |  | (1-0.014*tHb)*(HCO3--24.8+(1.43 x tHb +7.7)*(pH - 7.4)) |  |  | x
| HCO3- std | 25 | + | 0.78 | x | BE(B) | + | 0.002 | x | tHB | x | (O2Hb - 100) | mmol/L
No Co-Ox | HCO3- std | 25 | + | 0.78 | x | BE(B) | + | 0.002 | x | tHB | x | (sO2[c] - 100) | mmol/L
No Co-Ox | HCO3- (c) Actual Calculation |  |  | (pH | + | log(pCO2) | - | 7.608)
|  | 10 | ^
| TCO2 | HCO3- | + | 0.0307 | x | pCO2
| BE(B) Rolf Zander | (1 | - | 0.0143 | x | tHb) | x | (0.034 | x | pCO2 | x | 10^(pH-6.1) | - | 24.26) | + | (9.5 | + | 1.63 | x | tHb) | x | (pH-7.4)) | - | 0.2 | x | tHb | x | (100 | - | SAT) | / | 100
| SAT | 100 | / | [1 | + | (23400 | / | [pO2pp^3 | + | 150 | x | pO2pp])]
| pO2pp | pO2 | x | e^(C+0.003 x BE(B)) - 2.2) x (7.4 - pH)
| C | (pO2/26.7) | 0.184
| X | (1-0.014xtHb) | x | [HCO3- - 24 + (1.63xtHb+9.5) | x | (pH-7.4)
| tHb |  |  | 0.34*Hct
| Ca++ (7.4) | Ca++ | x | 10^[-0.178x(7.4-pH)}
| Anion GAP (AG) | (Na+ + K+) | - | (Cl- + HCO3-(c))
| P/F Ratio | paO2 | - | FiO2
| pAO2 | FiO2 | x | (BP-47) | - | 1.25 | x | paCO2(T)
| CaO2 | 0.0139 | x | tHb | x | O2Hb | + | 0.0031 | x | paO2(T)
| O2Ct | 0.0139 | x | tHb | x | O2Hb | + | 0.00314 | x | paO2(T)
sO2(c)
| CvO2 | 0.0139 | x | tHb | x | O2Hb | + | 0.0031 | x | pvO2(T)
| P50 mmHg | 10 | ^-(Q/2.7)
| Q | Log[R/ | (100-R)] | - | 2.7 | x | log | (pvO2(t))
| R | O2Hb | or | sO2
| O2Hb or sO2 | range | 30-75
| O2cap | (thb | - | tHb x (100-O2Hb-HHb) | x | 1.39
| A-aDO2 | pAO2 | - | paO2(T)
| paO2/pAO2 | paO2(T)/pAO2
| RI | A-aDO2 | / | paO2(T)
| CcO2 | (1.39 | x | tHb | x | alpha) | + | 0.00314 | x | pAO2
| alpha | (1-COHb/100) | - | C
| C | 0 | if pAO2 | >150
|  | 0.01 |  | >125 | - | <=150
|  | 0.02 |  | <=125
| a-vDO2 | CaO2 | - | CvO2
| Qsp/Qt(est) | 100 | x | (CcO2-CaO2) | / | (CcO2-CvO2)
| OI | (FiO2 | x | MAP) | / | paO2
| mOsm | (2xNa+) | + | Glu
| Hct (c) |  |  | 3.0*Hb
Temperature Correction
| pH(T) |  |  | pH+(T-37)*(0.0147+0.0065*(7.4-pH))
| pCO2(T) |  |  | pCO2*10^[0.019 x (T-37)]
| pO2(T) |  |  | pO2*10^[K x (T-37)]

## Result
Measured | [6 empty] | Corrected () |  |  | Co-Oximetry | [5 empty] | Derived
pH |  |  |  | pCO2 | pO2 | Hct | pH (T) | pCO2 (T) | pO2 (T) | tHb | O2Hb | COHb | MetHb | HHb | sO2 | TCO2 | Beecf | BE(B) | Ca++ (7.4) | AG | P/F Ratio | pAO2 | CaO2 | O2cap | O2ct | sO2(c) | HCO3-(c) | HCO3-std | A-aDO2 | paO2/pAO2 | RI | CcO2 | Qsp/Qt(est) | OI | Temp | FiO2 | BP
|  |  |  | mmHg | mmHg | % |  | mmHg | mmHg | g/dL | % | % | % | % | % | mmol/L | mmol/L | mmol/L | mmol/L | mmol/L | mmHg | mmHg | mmHg | mL/dL | mL/dL | % | mmol/L | mmol/L | mmHg |  |  | mL/dL | % | % | OC | % | mmHg
No | Nama RS / RSUD / RSUP | Operator | pH | pCO2 | pO2 | Hct | pH (T) | pCO2 (T) | pO2 (T) | tHb | O2Hb | COHb | MetHb | HHb | sO2 | TCO2 | Beecf | BE(B) | Ca++ (7.4) | AG | P/F Ratio | pAO2 | CaO2 | O2cap | O2ct | sO2(c) | HCO3-(c) | HCO3-std | A-aDO2 | paO2/pAO2 | RI | CcO2 | Qsp/Qt(est) | OI | Temp | FiO2 | BP
1 |  |  | 7.5 | 53 | 162 |  | 7.5 | 53 | 162 | 12.9 | 97 | 0.9 | 1.2 | 0.9 | 99.1 | 42.3511 |  | 20.159 | [8 empty] | 40.724 | [7 empty] | 37
2 | 1124041567 |  | 7.449 | 37.2 | 190.6 |  | 7.4404 | 36.2362 | 47.8766 | [6 empty] | 29.6175 |  | 4.4557 | [8 empty] | 28.4755 | [7 empty] | 36.4 | 80
[18 empty] | (1 | - | 0.0143 | x | tHb) | x | (0.034 | x | pCO2 | x | 10^(pH-6.1) | - | 24.26) | + | (9.5 | + | 1.63 | x | tHb) | x | (pH-7.4)) | - | 0.2 | x | tHb | x | (100 | - | SAT) | / | 100

## Cal
Measured | [6 empty] | Corrected () |  |  | Co-Oximetry | [5 empty] | Derived | [23 empty] | https://patologiklinik.com/2012/01/12/cara-mudah-membaca-analisa-gas-darah/
pH |  |  |  | pCO2 | pO2 | Hct | pH (T) | pCO2 (T) | pO2 (T) | tHb | O2Hb | COHb | MetHb | HHb | sO2 | TCO2 | Beecf | BE(B) | Ca++ (7.4) | AG | P/F Ratio | pAO2 | CaO2 | O2cap | O2ct | sO2(c) | HCO3-(c) | HCO3-std | A-aDO2 | paO2/pAO2 | RI | CcO2 | Qsp/Qt(est) | OI | Temp |  |  | FiO2 | BP | http://www-users.med.cornell.edu/~spon/picu/calc
|  |  |  | mmHg | mmHg | % |  | mmHg | mmHg | g/dL | % | % | % | % | % | mmol/L | mmol/L | mmol/L | mmol/L | mmol/L | mmHg | mmHg | mmHg | mL/dL | mL/dL | % | mmol/L | mmol/L | mmHg |  |  | mL/dL | % | % | OC |  |  | % | mmHg
No | Nama RS / RSUD / RSUP | Operator | pH | pCO2 | pO2 | Hct | pH (T) | pCO2 (T) | pO2 (T) | tHb | O2Hb | COHb | MetHb | HHb | sO2 | TCO2 | Beecf | BE(B) | Ca++ (7.4) | AG | P/F Ratio | pAO2 | CaO2 | O2cap | O2ct | sO2(c) | HCO3-(c) | HCO3-std | A-aDO2 | paO2/pAO2 | RI | CcO2 | Qsp/Qt(est) | OI | Temp | C | K | FiO2 | BP
1 |  |  | 7.15 | 69 | 98 | 0 | 7.16 | 68 | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 |  | 0
2 |  |  | 7.15 | 69 | 98 |  | 7.15 | 69 | [26 empty] | 37 | 1.2703 |  | 21 | 760
|  |  | 7.15 | 69 | 98 |  | 7.16 | 68 | 95 | 15.6 | 95.7 | 1.1 | 1 | 2.3 | 97.7 | 26.1 | -4.9 | -6.5 |  |  | 102 | 599 | 21 | 21.2 | 21 | 95.3 | 24 | 19.8 | 504 | 0.16 | 5.3 | 23.3 | 39.7
|  |  | 7.181 | 54.7 | 73.8 | 42.8 |  |  |  | 36.5 | [5 empty] | 21.4 | -9 | -7.8 |  |  |  |  | 21.4 |  | 18.9 |  | 19.7 | 17.2 | 555.4 |  | 7.52 |  |  |  | 36.5 |  |  | 96 | 760
| pH(T) | pH | + | (T-37) | x | [-0.0147 + 0.0065 x (7.4-pH)} |  | pH(T2) = pH + (temp - 37) * (-0.0147 + 0.0065 * (7.4 - pH))
| pCO2(T) | pCO2 | x | 10^[0.019 x (T-37)]
| pO2(T) | pO2 | x | 10^[K x (T-37)]

## Ar_Male_S
Patient
Lab Number |  | <<Lab Number>>
Medical Record |  | <<Medical Record>>
Name |  | <<NAME>>
Gender / Age |  | <<GENDER>> / <<AGE>>
Analyzer
Type |  | GEM® Premier 5000
Serial Number |  | <<Serial Number>>
Software Version |  | <<Software Version>>
Results
|  |  | Crit, | Reference |  | Crit.
|  |  | Low | Low | High | High
Corrected | (<<>> OC)
pH(T) |  |  | 7.2 | 7.35 | 7.45 | 7.6
pCO2(T) |  | mmHg | 20 | 35 | 48 | 70 | pH (Ven0us) | 7.32 | 7.43
pO2(T) |  | mmHg | 40 | 83 | 106
Measured | (37.0 OC) | [5 empty] | pCO2 (Female) | 32 | 45 | mmHg
pH |  |  | 7.2 | 7.35 | 7.45 | 7.6 | pCO2 (Venous) | 26 | 52 | mmHg
pCO2 |  | mmHg | 20 | 35 | 48 | 70
pO2 |  | mmHg | 40 | 83 | 106
Na+ |  | mmol/L | 120 | 136 | 145 | 160
K+ |  | mmol/L | 2.8 | 3.5 | 5.1 | 7.8
Cl- |  | mmol/L | 80 | 98 | 107 | 120
Ca++ |  | mmol/L | 0.75 | 1.15 | 1.33 | 1.6
Glu |  | mg/dL | 40 | 65 | 95 | 450 | Hct (Female) | 37 | 47 | %
Lac |  | mmol/L |  | 0.36 | 0.75 | 3.4
Hct |  | % |  | 40 | 50
Co-Oximetry
tHb |  | g/dL | 7 | 12.6 | 17.4 | 20
O2Hb |  | % |  | 90 | 95
COHb |  | % |  |  | 10
MetHb |  | % |  |  | 1.5
HHb |  | % |  | 1 | 5
sO2 |  | % |  | 94 | 98 |  | Lac (Vernous) | 0.56 | 1.39 | mmol/L
Derived
TCO2 |  | mmol/L |  | 19 | 24 |  | tHb (Female) | 11.7 | 16.1 | g/dL
Beecf |  | mmol/L
BE(B) |  | mmol/L |  | -2 | 3
Ca++(7.4) |  | mmol/L |  |  |  |  | COHb (Smokers) |  | 10 | %
AG |  | mmol/L |  | 10 | 20
P/F Ratio |  | mmHg
pAO2 |  | mmHg
CaO2 |  | mmHg
O2cap |  | mL/dL |  |  |  |  | TCO2 (Venous) | 22 | 26 | mmol/L
O2ct |  | Ml/dL
sO2(c) |  | %
HCO3-(c) |  | mmol/L | 10 | 21 | 28 | 40 | HCO3- (Venous) | 22 | 29 | mmol/L
HCO3-std |  | mmol/L
A-aDO2 |  | mmHg
paO2/pAO2
RI
CcO2 |  | mL/dL
Qsp/Qt(est) |  | %
OI |  | %
Entered
Temp |  | OC
FiO2 |  | %
BP |  | mmHg

## Ar_Male_NS
Patient
Lab Number |  | <<Lab Number>>
Medical Record |  | <<Medical Record>>
Name |  | <<NAME>>
Gender / Age |  | <<GENDER>> / <<AGE>>
Analyzer
Type |  | GEM® Premier 5000
Serial Number |  | <<Serial Number>>
Software Version |  | <<Software Version>>
Results
|  |  | Crit, | Reference |  | Crit.
|  |  | Low | Low | High | High
Corrected | (<<>> OC)
pH(T) |  |  | 7.2 | 7.35 | 7.45 | 7.6
pCO2(T) |  | mmHg | 20 | 35 | 48 | 70 | pH (Ven0us) | 7.32 | 7.43
pO2(T) |  | mmHg | 40 | 83 | 106
Measured | (37.0 OC) | [5 empty] | pCO2 (Female) | 32 | 45 | mmHg
pH |  |  | 7.2 | 7.35 | 7.45 | 7.6 | pCO2 (Venous) | 26 | 52 | mmHg
pCO2 |  | mmHg | 20 | 35 | 48 | 70
pO2 |  | mmHg | 40 | 83 | 106
Na+ |  | mmol/L | 120 | 136 | 145 | 160
K+ |  | mmol/L | 2.8 | 3.5 | 5.1 | 7.8
Cl- |  | mmol/L | 80 | 98 | 107 | 120
Ca++ |  | mmol/L | 0.75 | 1.15 | 1.33 | 1.6
Glu |  | mg/dL | 40 | 65 | 95 | 450 | Hct (Female) | 37 | 47 | %
Lac |  | mmol/L |  | 0.36 | 0.75 | 3.4
Hct |  | % |  | 40 | 50
Co-Oximetry
tHb |  | g/dL | 7 | 12.6 | 17.4 | 20
O2Hb |  | % |  | 90 | 95
COHb |  | % |  |  | 3
MetHb |  | % |  |  | 1.5
HHb |  | % |  | 1 | 5
sO2 |  | % |  | 94 | 98 |  | Lac (Vernous) | 0.56 | 1.39 | mmol/L
Derived
TCO2 |  | mmol/L |  | 19 | 24 |  | tHb (Female) | 11.7 | 16.1 | g/dL
Beecf |  | mmol/L
BE(B) |  | mmol/L |  | -2 | 3
Ca++(7.4) |  | mmol/L
AG |  | mmol/L |  | 10 | 20
P/F Ratio |  | mmHg
pAO2 |  | mmHg
CaO2 |  | mmHg
O2cap |  | mL/dL |  |  |  |  | TCO2 (Venous) | 22 | 26 | mmol/L
O2ct |  | Ml/dL
sO2(c) |  | %
HCO3-(c) |  | mmol/L | 10 | 21 | 28 | 40 | HCO3- (Venous) | 22 | 29 | mmol/L
HCO3-std |  | mmol/L
A-aDO2 |  | mmHg
paO2/pAO2
RI
CcO2 |  | mL/dL
Qsp/Qt(est) |  | %
OI |  | %
Entered
Temp |  | OC
FiO2 |  | %
BP |  | mmHg

## Male_S
Patient | [6 empty] | Analyzer
Lab Number |  | <<Lab Number>> |  |  |  |  | Type |  | GEM® Premier 5000 |  |  |  |  | No | Nama RS / RSUD / RSUP | pH | pCO2 | pO2 | Hct | tHb | O2Hb | COHb | MetHb | HHb | sO2 | Temp | FiO2 | BP
Medical Record |  | <<Medical Record>> |  |  |  |  | Serial Number |  | <<S N>> |  |  |  |  | 1 | Sehat Senantiasa | 7 | 30 | 100 | 40 | 13.3 | 95 | 2 | 0 | 3 | 98 | 36.4 | 21 | 760
Name |  | <<NAME>> |  |  |  |  | Software Version |  | <<S V>> |  |  |  |  | 2 | abadi
Gender / Age |  | <<Gender>> |  | / | <<Age>>
Results
|  |  | Crit, | Reference |  | Crit. |  |  |  | Crit, | Reference |  | Crit.
|  |  | Low | Low | High | High |  |  |  | Low | Low | High | High
Corrected | (<<Temp>> OC) | [5 empty] | Derived
pH(T) | 7.39 |  | 7.2 | 7.35 | 7.45 | 7.6 | TCO2 | 29.8939 | mmol/L |  | 19 | 24 |  |  | TCO2 = HCO3- + 0.0307 x pCO2
pCO2(T) | 47 | mmHg | 20 | 35 | 48 | 70 | Beecf | 3.489 | mmol/L | [5 empty] | BE(ecf) = HCO3- - 24.8 + 16.2 x (pH - 7.4)
pO2(T) | 52 | mmHg | 40 | 83 | 108 |  | BE(B) | 3.2848 | mmol/L |  | -2 | 3 |  |  | BE(B) mmol/L = (1 - 0.014 x tHb) x [HCO3- - 24.8 + (1.43 x tHb + 7.7) x (pH - 7.4)]
Measured | (37.0 OC) | [5 empty] | Ca++(7.4) | 1.1254 | mmol/L | [5 empty] | Ca++(7.4) = Ca++ * 10(-0.178 * (7.4 - PH))
pH | 7.39 |  | 7.2 | 7.35 | 7.45 | 7.6 | AG | 71.751 | mmol/L |  | 10 | 20 |  |  | AG = (Na++ + K+) - (Cl- + HCO3-)
pCO2 | 47 | mmHg | 20 | 35 | 48 | 70 | P/F Ratio | 0 | mmHg
pO2 | 52 | mmHg | 40 | 83 | 106 |  | pAO2 | 149.1425 | mmHg | [5 empty] | pAO2 mmHg = FIO2 x (BP – 47) – 1.25 x paCO2(T) |  |  | Alveoral gas ecuation | PAO2 = ( FiO2 * (Patmos - PH2O)) - (PaCO2 / RQ) |  |  |  |  | PAO2 = ( FiO2 * (760 - 47)) - (PaCO2 / 0.8)
Na+ | 140 | mmol/L | 120 | 136 | 145 | 160 | CaO2 | 9.7112 | mmHg | [5 empty] | CaO2 mL/dL = 0.0139 x tHb x O2Hb + 0.0031 x paO2(T)
K+ | 3.3 | mmol/L | 2.8 | 3.5 | 5.1 | 7.8 | O2cap | -8.8543 | mL/dL | [5 empty] | O2cap mL/dL = (tHb - tHb x (100-O2Hb -HHb)) x 1.39
Cl- | 100 | mmol/L | 80 | 98 | 107 | 120 | O2ct | 9.7133 | Ml/dL | [5 empty] | O2ct (mL/dL) = (0.0139 x tHb x O2Hb) + 0.00314 x pO2(T) |  |  | Oxygen Content | CaO2 = (1.36 * Hgb *SaO2+ 0.0031 * PaO2)/100
Ca++ | 1.13 | mmol/L | 0.75 | 1.15 | 1.33 | 1.6 | sO2(c) | 86.3801 | % | [5 empty] | sO2(c) = (23400*(pO23 + 150*pO2)^-1 + 1)^-1)*100
Glu | 222 | mg/dL | 40 | 65 | 95 | 450 | HCO3-(c) | 28.451 | mmol/L | 10 | 21 | 28 | 40 |  | HCO3-(c) = 10(pH + log(pCO2) – 7.608)
Lac | 2.2 | mmol/L |  | 0.36 | 0.75 | 3.4 | HCO3-std | 27.1162 | mmol/L | [5 empty] | HCO3-std = 25 + 0.78 x BE(B) + 0.002 x tHb x (O2Hb -100)
Hct | 29 | % |  | 40 | 50 |  | A-aDO2 | 102.1425 | mmHg | [5 empty] | A-aDO2 mmHg = pAO2 – paO2 (T) |  |  |  | A-a gradient = PAO2 - PaO2
Co-Oximetry | [6 empty] | paO2/pAO2 | 0.34866
tHb | 9.1 | g/dL | 7 | 12.6 | 17.4 | 20 | RI | 1.9643 | [6 empty] | RI = A-aDO2 / paO2(T)
O2Hb | 75.5 | % |  | 90 | 95 |  | CcO2 | 12.8517 | mL/dL | [5 empty] | CcO2 mL/dL= (1.39 x tHb x alpha) +0.00314 x pAO2
COHb | 1.1 | % |  |  | 10 |  | Qsp/Qt(est) | 47.293 | % | [5 empty] | Qsp/Qt (est) % = 100 x (CcO2 – CaO2) / (3.5 + CcO2 – CaO2)
MetHb | 0.5 | % |  |  | 1.5 |  | OI | 0 | % | [5 empty] | OI = (FiO2 x MAP) / paO2 |  |  | Oxygen Index | OI = (FiO2 * MAP)/PaO2
HHb | 22.8 | % |  | 1 | 5 |  | Hct(c) | 27.3 | % | [5 empty] | Hct(c)= 3.0 * tHb
sO2 | 76.8057 | % |  | 94 | 98 |  | tHb(c) | 9.86 | g/dL | [5 empty] | tHb(c) = 0.34 * Hct
Entered | [14 empty] | sO2(oximetri) = O2Hb/(O2Hb+HHb)*100 |  |  |  | SO2 = (23,400 * (pO23 + 150 * pO2)-1 + 1)-1
Temp | 37 | OC
FiO2 | 21 | %
BP | 760 | mmHg
Help Formula
K | 1.1305
e | 2.718
pO2pp | 0.18024
C | 0.01 | [13 empty] | C = 0, if pAO2 is > 150 C = 0.01, if pAO2 is > 125, but ≤150 C = 0.02, if pAO2 is ≤ 125
alpha | 0.979 | [13 empty] | alpha = (1 - COHb/100) – C